## Slide 1: Quality Management System for Software Development
Gijsbert Bol

## Slide 2: Introduction (1)
UMCU computer science team
  4 core members
  Clinical software
    Delineation software
    Dicomserver
    Workflow software
  Pre-clinical software
    MRLTP
    Tracking/gating software
    Autosegmentation
    MR patches
  Research software

## Slide 3: Introduction (2)
Archiving, versioning and documentation
  Source code repositories
    Git
  Test reports
Responsibility
  Clinical physics
First-in-man MRLinac treatment (2017)
  Relied heavily on in-house software
  Technical file required
    IGZ and internal METC
  Regulations? Norms?
  2020!!!

## Slide 4: Regulations and Norms
MDD 93-42EEG Medical Devices Directive
MDR-2017/745 Medical Device Regulations

## Slide 5: Regulations and Norms
MDD 93-42EEG Medical Devices Directive
MDR-2017/745 Medical Device Regulations
ISO 13485 Medical devices – Quality management systems - Requirements for regulatory purposes
ISO 14971 Medical devices - Application of risk management to medical devices
ISO/IEC 62304 Medical device software Software life cycle processes

## Slide 6: Regulations and Norms

## Slide 7: Legal framework
Medical Device Regulation 2017 (MDR)
MDR aims:
  Balances patient care and small / medium-sized enterprice interests
  High standard regarding quality and safety regarding MDs
  Data generated by MDs are reliable and robust
  Protection of subjects participating in clinical investigations
  Applies to the whole EU
  Adopted from May 26th, 2020

## Slide 8: Medical device definition
Any instrument, apparatus, appliance, software, implant, reagent, material or other article intended by the manufacturer to be used, alone or in combination, for human beings for one or more of the following specific medical purposes:
  Diagnosis, prevention, monitoring, prediction, prognosis, treatment, compensation or alleviation of disease, injury or disability
  Investigate, replacement or modification of anatomy or a physiological or pathological process

## Slide 9: Software can be a medical device
MDR definition:
  Software shall also be deemed to be an active device
Active device:
  Any device, the operation of which depends on a source of energy and which acts by changing the density of or converting that energy
  Software can be regarded as a medical device with a power cord

## Slide 10: Classification (ISO 13485)
Governed by the intended purpose of the devices
Software, which drives a device or influences the use of a device, shall fall within the same class as the device
If the software is independent of any other device, it shall be classified in its own right
Classification determines the extend in which a appropriate authority assess the development process and the documentation
  Notified body (CE)
  IGZ

## Slide 11: Classification (ISO 13485)
Provides information to take decisions with diagnosis or therapeutic purposes
  class IIa
Except:
  Death or an irreversible deterioration of a person's state of health
    class III
  Serious deterioration of a person's state of health or a surgical intervention
    class IIb

## Slide 12: Classification (ISO 13485)
Monitor physiological processes
  class IIa
Monitor vital physiological processes, variation of the parameter may result in immediate danger to patient
  class IIb
All other:
  class I
MEDDEV 2.4/1 rev. 9
CE certification by a notified body needed:
  Class IIa, IIb, III

## Slide 13: Classification (IEC 62304)

## Slide 14: Hospitals (1)
Manufacturing, modifying and using devices in-house
On a non-industrial scale
Specific needs of target patient groups which cannot be met at the appropriate level of performance by an equivalent device available on the market
Within the an hospital setting, it is appropriate to provide that certain rules of MDR should not apply, since the aims of this MDR should still be met in a proportionate manner.

## Slide 15: Hospitals (2)
Software is not transferred to another legal entity
Software is developed and used under a appropriate quality management system (QMS)
Prove and document that other commerical software is not available for the target patient group’s specific needs
Provides information to the authorities:
  Justifying the development, modification and use of the software
  Understanding the development process, design and indended use of the sofware
  Asserting the general safety and performance requirements

## Slide 16: Hospitals (3)
Publicly declare that the software meets the general safety and performance requirements
Hospital reviews experience gained
  Takes all necessary corrective actions

## Slide 17: General software development requirements (1)
Risk management plan
  Identify foreseeable hazards
  Estimate risks during intended use and foreseeable misuse
  Eliminate or control the estimated risks
  Evaluate the impact and amend control measures

## Slide 18: General software development requirements (2)
Design and development
  Software development life cycle
    V-model:

## Slide 19: General software development requirements (3)
Technical documentation
  Specification, requirements, design, development, safety
  Verification and validation
    Verification: Are we building the system right?
    Validation: Are we building the right system?
  Pre-clinical and clinical data

## Slide 20: Quality Management System (1)
Provides:
  QMS user roles: task / responsibility / competence
  Procedures
    Risk managment procedures
    Quality management procedures
    Development procedures
  Forms
Stores for each software project:
  Risk analyis data
  Verification and validation data
  Test data
Produces for each software project:
  Technical file
    Notified body (CE)
    IGZ

## Slide 21: Quality Management System (2)

## Slide 22: Quality Management System (3)
Norms:
  ISO 13485 and ISO 14971
    General QMS for medical devices
  ISO 62304
    QMS for medical software
    Software development life cycle processes
QMS certification:
  A QMS can be certified under ISO 13485 and ISO 14971 only

## Slide 23: ISO 62304 overview

## Slide 24: UMCU implementation
Git
Latex
PDF publishing
  Electronic signing

## Slide 25: UMCU implementation
Git server
    gitolite
Build (unit) test system
    maven
Document Management System
  Review/approve
    seeddms
Issue tracker
Git
Latex
PDF publishing
  Electronic signing

## Slide 26: Involvement
Requirements
Risk analysis
Implementation
Testing

## Slide 27: FAQ
Is configuring == programming?
  No!
  Within intended use of manufacturer
Can scripts be medical devices?
  Yes!
  Medical device definition in MDR
  Independent of programming language
  Within intended use of manufacturer
Can Excel sheets be medical devices?
  Yes!
  Medical device definition in MDR
Can programs / scripts / Excel sheets be shared?
  Only with a CE mark

## Slide 28: Links
Praktijkgids Medische Informatietechnologie (2018)
  https://mtintegraal.nl/specials/1/speciale-uitgave-praktijkgids-medische-informatietechnologie
“HET ONTWIKKELEN VAN MEDISCHE SOFTWARE IS EEN VAK WAARBIJ JE ONDER ANDERE ERVAREN ONTWIKKELAARS, PROJECTLEIDERS EN ARCHITECTEN NODIG HEBT. HET IS STERK AF TE RADEN ZELF TE ONTWIKKELEN ALSDEZE ERVARING EN KENNIS NIET IN HUIS IS.”

## Slide 29: Conclusions
Adhere to norms
  ISO 62304, ISO 13485 and ISO 14971
Implement a software QMS
  Procedures, roles, records
  Programs, scripts and Excel sheets
Setup team for each software project
  Developers (risk analysis, implementation, testing)
  Clinical physicists (requirements, risk analysis, testing)
  End users (requirements, risk analysis, testing, bug reports)