## Slide 1: RT-24
Workgroup radiotherapy Belgium 2023-05-25
BelDART+External beam audit
Chair: Bert bakelandt

## Slide 2: BELdART + external dosimetry audit audit
Scope of the meeting:
discuss BELdART results + give feedback to the steering committee
discuss mandatory external dosimetry audits to come to a conclusion at the RT-24 live meeting

## Slide 3: BELdART dosimetry audit: 1 example

## Slide 4
BELdART dosimetry audit: 1 example

## Slide 5
BELdART dosimetry audit: 1 example

## Slide 6: BELdART dosimetry audit

## Slide 7: BELdART dosimetry audit

## Slide 8: (BELdART ) dosimetry audit

## Slide 9: (BELdART ) dosimetry audit

## Slide 10: Dosimetry with alanine
Irradiation of the amino acid alanine produces stable free radicals:
The concentration of the radicals is proportional to the absorbed dose and can be measured by Electron Paramagnetic Resonance (EPR) spectroscopy
M. Anton, Applied Radiation and Isotopes 62 (2005), 779–795

## Slide 11
Results IAEA blind irradiations
Co-60 beams
Co-60 beams

## Slide 12: Results IAEA blind irradiations
Linac based (using BELdART holder)

## Slide 13: Results IAEA blind irradiation tests
Results other DAN’s
https://dosimetry-audit-networks.iaea.org/Home/AuditQA

## Slide 14: Uncertainties obtained from PTB
AD | 
25 | : dose of basis pellets in Gy
4 | : number of base pellets (4 or 8)
4 | : dose of test pellet in Gy
4 | : number of test pellets per dose
0.3% | : intrabatch homogeneity, single pellet
0.020 | : individual background Am (single pellet) in Gy
0.050 | : repeatability Am (single pellet) in Gy
primary standard | 0.30%
test probes | 1.10%
reproducibility of irradiation | 0.50%
irradiation temperature | 0.10%
intrabatch homogeneity | 0.15%
average mass of n probes | 0.05%
ESR amplitude | 0.67%
fading correction | 0.09%
environment irradiation different from cal. | 0.50%
radiation quality different from Co-60 or Cs-137 | 0.45%
base function | 0.25%
reproducibility of irradiation | 0.05%
irradiation temperature | 0.05%
intrabatch homogeneity | 0.15%
average mass of n probes | 0.05%
ESR amplitude | 0.11%
"systematic" part (batch hl02/kapL et al.) | 0.15%
Relative combined standard uncertainty u(N) / N = | 1.164%

## Slide 15: BELdART-1 => other BELdARTs
Dose verification of basic and dynamic RT techniques in Belgium
4 objectives:
  Basic dosimetric check of MVX beams
    Subset of BELdART tests
    Mailing audit based on BELdART-1 protocols using dedicated holder
  Dynamic RT (IMRT, tomo, RA, VMAT, SRS, SBRT, ...)
    Mailing audit: phantom preloaded with alanine dosimeters & film
    Scientific committee in collaboration with BHPA
  MeV Electron audits
  Benchmarking with IAEA for the alanine

## Slide 16: Basic audit: standard linacs
Test 1 and 2: reference beam output
  8 Gy at the usual at 8cm, 10cm and at 20cm in SSD or SAD technique as is used for clinic with a vertical beam set up and a 10 x 10 cm2 field.
Test 2b: output wedged field (if used clinically, on special request)
  8 Gy should be delivered at 8 cm depth in water. Half the dose (4 Gy) should be administered for a wedge orientation corresponding to a 0° and 180° collimator rotation respectively
Test 3: beam output for different SSD
  8 Gy at 8cm for SSD 85 with a vertical beam set up and a 10 x 10 cm2 field.
Test 4: output small Field > 2x2cm2
  Solid phantom: 8 Gy at 5cm water equivalent
  Future: back to water!

## Slide 17: Electron beams
Use of spacers to reach Z ref

## Slide 18: Film dosimetry: EBT3 & EBT-XD
Calibration 0, 1, 2, 4, 8, 16, (32) Gy
  against alanine for 8Gy and 16Gy
”One scan dosimetry” used in the hospitals to free us from time and environmental constraints
Intercomparison with IAEA in progress

## Slide 19: E2E tests: IMRT/VMAT
80%
50%
80%
50%

## Slide 20
STEEV phantom
Use 2 pieces of film
  1 film in combination with alanine 1 film without alanine
  Films are placed in axial plane
  All the films are read together
BELdART-SRS
Notes: -water equivalent (no tumour material)
- Cushion ( as air equivalent TRANSPARANT as possible to not influence the dose to the head)

## Slide 21
Beldart-SBRT
BELdART - Nov 21, 2022
Alanine pellet
Film location on top of lungs
Film location through tumour
Notes: 21

## Slide 22: Some statistics since 2016 …
Number of audits: 79
Number of beams: 204 photons and electrons
Number of electron audit: 32
Number of basic tests only audits: 15
Number of photon VMAT audits: 79
Number of photon SRS audits: 46
Number of photon SBRT audits: 28

## Slide 23: Basic tests
Chart: Distribution of used energies in basic tests
| 6 MV | 6 MV FFF | 7 MV FFF | 10 MV | 10 MV FFF | 15 MV | 18 MV
 | 41.6667 | 22.0238 | 0.59524 | 14.2857 | 4.7619 | 12.5 | 4.1667
Chart: distribution of FF/FFF audits
| FF beams | FFF beams
 | 72.0238 | 27.9762

## Slide 24: End2end: VMAT
Chart: End-to-End alanine-EPR results beams
3%
-3%
Chart: Distribution used algorithms
| AAA | Acuros Dm | CCC | Convolution Superposition | Monte Carlo | VoLO (GPU Collapsed Cone Superposition)
 | 32.9114 | 10.1266 | 18.9873 | 1.2658 | 34.1772 | 2.5316
Chart: Gamma analysis 3%/3mm beams
| nr. of beams >98% | nr. of beams 98%-95% | nr. of beams 95%-90% | nr. of beams <90%
 | 68.3544 | 20.2532 | 6.3291 | 5.0633

## Slide 25: End2end: SRS
Chart: Distribution used algorithms
| AAA | Acuros Dm | Adaptive CC | CCC | Monte Carlo | PBC | Ray Tracing
 | 17.3913 | 13.0435 | 2.1739 | 15.2174 | 39.1304 | 10.8696 | 2.1739
Chart: SRS alanine/EPR results
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46
SRS | 1.284 | -0.56051 | -1.4116 | 3.7935 | 1.7915 | 3.7273 | 2.3383 | 2.5007 | 2.8127 | -0.05357 | 2.1681 | 1.7939 | 3.1015 | 1.5183 | -2.109 | -1.6996 | 1.9554 | -0.0074406 | 0.44382 | -0.68288 | 2.8595 | 4.4875 | -0.23387 | -0.7289 | 1.8135 | 1.3329 | 1.9293 | 1.7298 | 2.0284 | 1.2985 | 1.0104 | 0.86091 | 0.56853 | 0.84384 | 0.13177 | 1.6017 | 2.4864 | -6.0639 | -6.7074 | -6.2922 | -6.471 | -0.62776 | -0.503 | -0.044826 | -0.057181 | 3.8393
Beam output | 0.07479 | 0.2061 | 0.36519 | -1.5511 | 0.070494 | -0.31341 | 1.0095 | -0.77291 | 1.7389 | 0.28827 | 0.8805 | -1.6444 | 0.27269 | 0.23078 | -1.2548 | -0.44515 | -1.1369 | -0.11557 | -0.44816 | -0.97225 | -0.87123 | -1.4953 | -2.0795 | -0.021402 | -0.97376 | -1.8696 | -2.042 | -1.3125 | -1.9031 | -1.732 | -1.7144 | -0.99207 | -1.3083 | -1.2104 | -0.97224 | 0.36126 | -0.65093 | 0.53806 | 1.9018 | -0.60481 | -1.2172 | -2.2087 | -1.168 | -1.1154
Chart: Gamma analysis 5%/1mm 50% LD threshold large lesion
| % of beams >98% | % of beams 98%-95% | % of beams 95%-90% | % of beams <90%
Large lesion | 86.9565 | 6.5217 | 0 | 6.5217
film

## Slide 26: End2end: SBRT
Chart: Distribution used algorithms
| AAA | Acuros Dm | Adaptive CC | CCC | Monte Carlo | PBC | Ray Tracing
 | 21.4286 | 35.7143 | 0 | 14.2857 | 28.5714 | 0 | 0
w
Chart: Gamma analysis through target 5%/1mm
| % of beams >98% | % of beams 98%-95% | % of beams 95%-90% | % of beams <90%
Through target(5%/1mm) | 82.1429 | 10.7143 | 7.1429 | 0
Chart: SBRT alanine/EPR results
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28
SBRT | -0.78403 | 1.3114 | 2.2578 | 0.81624 | 0.41472 | 0.96873 | 3.6762 | 3.8 | 1.8471 | 2.3447 | 2.883 | 2.3244 | 1.909 | 2.6162 | 4.6539 | 4.1154 | 3.837 | 2.8257 | 4.0925 | 3.6221 | 2.9014 | 4.6794 | 5.727 | 3.5544 | 1.8841 | 2.3601 | 2.7778 | 1.3731
Beam output | -1.063 | -0.15117 | -0.44515 | -0.11263 | -2.1281 | 1.3163 | -0.97721 | -2.7548 | -0.058706 | -0.87123 | -1.4953 | -0.021402 | -0.97376 | -0.31095 | -0.6467 | -1.8696 | -2.042 | -1.3125 | -1.9031 | -1.732 | -1.7144 | -0.99207 | -1.3083 | -0.4136 | -0.21235 | -2.2061 | -2.2087 | -1.168
film

## Slide 27: Future work
New table-top spectrometer
  As backup for dosimetry (reference is ruby)
  For higher doses?
  Other materials?
Collaboration with CIBIO in France
Intercomparison planned with VSL in Catharina hospital in Eindhoven
Development of new phantoms
Breathing lung phantom (project with Maastro)
Expend the use of 3D printing
Brachytherapy!

## Slide 28: Future work
New MR-linacs?
  One test performed (no 3D printed supports to avoid air pockets)
  MC models should be developed?
Testing of the new type of gafchromic film from Ashland: EBT4
Lots of data Possible studies:
  Influence of the calculation algorithm?
  Include complexity? (Code from Barcelona)
  …
invesigate the effect of LET on the spectra.

## Slide 29: BELdART
Call with Brigitte Reniers:
Any problems with the audits? ‘Not a lot of problems (+5%), only 1 center problem with reference dose. Unclear if the center resolved the issue and reaudited else.
More then 150 beams measured
Outliers in SRS or SBRT could not be measured again due to reasons
BELdART gets IAEA audits: within 1%
There seems to be no problem with BELdART audit but this is not how several centers feel. We need to feel confident in the procedure…
We should signal this feeling towards BELdART -> steering committee

## Slide 30: Legislation

## Slide 31: Legislation

## Slide 32: Legislation
Important words:
External RT
+ 1 MeV
Before going clinically
External audit
Help and input from RT-24
Early ‘24

## Slide 33: Legislation
Important words:
External RT: brachy is out of the scope
+ 1 MeV: KV contact is out of the scope

## Slide 34: Legislation
What do we consider a minimum requirement for an external dosimetry audit?
Knowing that…….
timing is crucial in the process of commissionning a new machine
delay because of FANC/AFCN license issues
delay because of technical issues (infrastructure bunker/interlocks)
delay because of radiation protection
delay because of staffing problems, people getting ill,…
delay because of other priorities,….
delay because of ……..

## Slide 35: Legislation
What do we consider a minimum requirement for an external dosimetry audit?
Knowing that…….
we buy machines for different modalities
machines for 3D
machines for IMRT
machines for VMAT
machines for (IMRT/VMAT) SRS
machines for (IMRT/VMAT) SBRT
machines for photons, electrons, protons
machines with surface scanning
machines with CBCT imaging, CT imaging, kV-imaging
machines with MRI imaging

## Slide 36: Legislation
What do we consider a minimum requirement for an external dosimetry audit?
Do we need a simple reference field check, or do we feel comfortable if our complex treatment techniques are crosschecked?
Do we think a 10x10 reference field check is good practice if we will use the machine as a small field VMAT SRS treatment machine?
Maybe my machine is ready for conventional treatment (and pressure to go live as soon as possible) but it will take me some more months to get the complex SBRT treatment modelled in my TPS and validated.
Do we want to certify the external dosimetry audit or do I think it’s ok if my neighbouring RT department comes over to check my output?

## Slide 37: Legislation
Proposal to start the discussion:
The minimal requirement for an external dosimetry before going clinically is a basic output test done by a certified external audit organism.
If the machine is intended for a specific treatment it is good practice to order an independant audit within the first 3 months of the start of the specific treatment.

## Slide 38: Legislation
1.
2. Example and history of AZ Sint-Jan audit + re audit + finally Beldart on site with results which are satisfying, but loss of confidence in audit technique.
AZ St. Lucas had lots of problems:
Alanine was for several weeks at high temperature (35° C)
Measurements are rescaled to the output of the 10x10 reference field
Another centre came over to check the absolute dosimetry: was ok.
Hasselt experienced similar issues with basic audit test. They use Alanine in their centre and their own results are ok.
3. Proposal from Dirk Verellen to have an ionisation chamber measurement if the alanine measurement fails/is doubtful.
4. The basic test should be more robust - more professional.
5. Steering committee will put it on the agenda to discuss with BELdART
2.1 Legislation: important to keep the modalities in our own hands.
2.2 We need to be careful that the audit is feasable. In the same time we do think it's important to have a basic audit test for photons, protons, electrons before going clinically. There should be a transition period for specific techniques. We don't want specific techniques audits mandatory in the legislation. Lower the bar - but not to much.
2.3 The use of traffic light protocol to decide
2.4 Also beam matched machines need audit
2.5 Discussion on how electron audit could be performed. Only absolute dosimetry or some specific setup beams?
2.6 Can we audit ourselves as recognized MPE's or do we need a certified organism? The certified organism like BELdART itself is employing unrecognised employees. Do we list 'trained MPE's to audit?' In the past problems with that: the same MPE's have to go to audits all the time.
2.7 Concerning legislation: we should keep it simple, otherwise it will not be feasable to plan and accomplish the external audit