DRAFT version 05-2007: QC of conventional linear accelerators used for 3DCRT: criteria of acceptability

Document prepared by the workgroup of the BHPA.

Members: Marie-Thérèse Hoornaert, Odile Michel, Geert Pittomvils, Alex Rijnders, Aude Stembert, Michel Vandycke, Frederik Vanhoutte, Stefaan Vynckier, and Karen Haest (FANC).

General remarks:

This document presents criteria for acceptability with a proposal for tolerance and immediate action levels and verification frequencies. These tables still need to be translated into actual test procedures, where certain criteria can be globally verified in a single test (eg mechanical and optical isocenter). Such tests can be designed by the individual user, or found in the literature. Some examples will be presented on the BHPA website.

The criteria are presented for the standard use of a linac for 3DCRT. Certain criteria can however be adapted to the clinical practice of an individual institute/unit: items that are not used should not be tested (eg table position linear scales), while in other situations certain tolerances might need to be set stricter.

All tolerances expressed as a distance refer to the distance at the level of the isocenter of the accelerator.

No criteria are presented for specific accessories such as electronic portal imaging and verification system. The workgroup is aware that the need exists to include these systems in a comprehensive QA program, and strongly advises the users to do so, but we felt that it was out of the scope of the current group to prepare guidelines for this at this moment.

In the same manner no criteria are presented for patient immobilization and positioning systems, nor for more recent developments such as cone beam CT, gating and IMRT. Although these systems certainly demand a thorough QC program, we felt unable to present a global proposal applicable to all systems and would advise the users to develop a program adapted to the system(s) they are using.

Also QC guidelines for specific treatment techniques should be further developed.

The workgroup used the following documents as reference:

BHPA: draft document of the workgroup of 1995

Swiss society “QC of Medical Electron Accelerators”

CAPCA “Medical Linear Accelerators”

AAPM TG40

Any comments/remarks on this document are appreciated. Please mail to president@bhpa.be .

Table I.	Mechanical checks for conventional linear accelerators for 3DCRT

| PARAMETER | CRITERION | RECOMMENDED FREQUENCY | MINIMAL FREQUENCY | TOLERANCE LEVEL | IMMEDIATE ACTION LEVEL |
| --- | --- | --- | --- | --- | --- |
| Mechanical Axes |  |  |  |  |  |
| Mechanical isocenter Collimator rotation axis Gantry rotation axis Couch rotation axis | Accuracy Accuracy Accuracy | Monthly Monthly Monthly | Half-yearly Half-yearly Half-yearly | ±1 mm ±1 mm ±2 mm | ±2 mm ±2 mm ±3 mm |
| Angular scales (mechanical, analog and digital) Gantry rotation Collimator rotation at gantry 90° Collimator 0° at gantry 0°: field edge parallel to gantry rotation plane Gantry angle dependence of 2.b and 2.c | Accuracy Accuracy Accuracy Accuracy | Monthly Monthly Monthly Half-yearly | Half-yearly Half-yearly Half-yearly Yearly | ±0.5 ° ±0.5 ° ±0.5 ° ±1° | ±1° ±1° ±1° ±2° |
| Collimator jaws Opposing jaws parallel Jaw pairs perpendicular Gantry angle dependence of 3.a and 3.b | Accuracy Accuracy Accuracy | Monthly Monthly Half-yearly | Half-yearly Half-yearly Yearly | ±0.5 ° ±0.5 ° ±1.0 ° | ±1° ±1° ±2° |
| Collimator scales Jaw position < 10 x 10 cm² > 10 x 10 cm² Symmetric/asymmetric mode switching Non-divergent field match Maximum over/underdose at clinical SSD and depth FWHM over/underdose area at clinical SSD and depth Gantry angle dependence of 4.a and 4.c | Accuracy Accuracy Functionality Accuracy Accuracy Accuracy | Monthly Monthly Monthly Weekly Monthly Half-yearly | Half-yearly Half-yearly Half-yearly Monthly Half-Yearly Yearly | ±1 mm ±1% of field size (max. ±2 mm) Functional ±10% of local dose 2 mm idem | ±2 mm ±2% of field size (max. ±4 mm) Functional ±20% of local dose 4 mm idem |
| Multileaf Collimator Alignment with parallel jaw pair Alignment with perpendicular jaw pair Individual leaf position Field shape of representative test pattern Constancy of dosimetric parameters (leaf transmission, interleaf leakage) Gantry angle dependence of 5.a - 5.d | Accuracy Accuracy Accuracy Accuracy Constancy Accuracy | Monthly Monthly Monthly Monthly Half-yearly Half-yearly | Half-yearly Half-yearly Half-yearly Half-yearly Yearly Yearly | ±1° ±1° ±1 mm ±1 mm ±2% of local dose idem | ±2° ±2° ±2 mm ±2 mm ±4% of local dose idem |
| Couch axes Vertical axis parallel to beam axis over clinical range Maximum vertical displacement at isocenter under load Longitudinal axis parallel to in-plane over clinical range Lateral axis parallel to cross-plane over clinical range Couch top rotation | Accuracy Accuracy Accuracy Accuracy Accuracy | Monthly Half-yearly Monthly Monthly Half-Yearly | Half-yearly Yearly Half-yearly Half-yearly Yearly | ±2 mm ±5 mm ±1° ±1° ±1° | ±4 mm ±10 mm ±2° ±2° ±2° |
| Couch scales (mechanical, analog and digital) Linear scales over clinical range Angular scales over clinical range | Accuracy Accuracy | Half-yearly Half-yearly | Yearly Yearly | ±2 mm ±1° | ±4 mm ±2° |
| Optical and mechanical indicators |  |  |  |  |  |
| Laser alignment Laser visibility Isocenter Alignment along in-plane, cross-plane and beam axes over clinical range | Functionality Accuracy Accuracy | - Weekly Weekly | - Monthly Monthly | Functional ±1 mm ±1° / ±2 mm | Functional ±2 mm ±2° / ±4 mm |
| Optical reticule / crosswires Light intensity Isocenter coincidence at reference SSD In-plane/ Cross-plane coincidence at reference SSD Beam axis coincidence at clinical range SSD Gantry angle dependence of 9.a and 9.b | Functionality Accuracy Accuracy Accuracy Accuracy | - Weekly Weekly Monthly Monthly | - Monthly Monthly Half-yearly Half-yearly | Functional ±1 mm ±1° ±2 mm ±2 mm/ ±1° | Functional ±2 mm ±2° ±4 mm ±4 mm/ ±2° |
| Optical SSD indicator Light intensity Reference SSD Clinical range of SSD Gantry angle dependence of 10.b | Functionality Accuracy Accuracy Accuracy | - Weekly Monthly Monthly | - Monthly Half-yearly Half-yearly | Functional ±1 mm ±2 mm ±2 mm | Functional ±2 mm ±4 mm ±4 mm |
| Laser, optical and SSD visibility are supposed to be inherently checked by the operators/treatment technicians. Instructions should exist that whenever problems are experienced with the visibility of these systems a responsible person is notified and if needed corrective action is organized. |

| Accessories |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Wedge Detection of wedge type and direction Latching of wedge | Functionality Functionality | Monthly Monthly | Half-yearly Half-yearly | Functional Functional | Functional Functional |
| Tray Detection of tray coding Latching of tray Position | Functionality Functionality Constancy | Monthly Monthly Monthly | Half-yearly Half-yearly Half-yearly | Functional Functional ±2 mm | Functional Functional ±4 mm |
| Applicator Detection of applicator type and aperture insert Latching of applicator Position Collimator settings | Functionality Functionality Constancy Functionality | Monthly Monthly Monthly Monthly | Half-yearly Half-yearly Half-yearly Half-yearly | Functional Functional ±2 mm Functional | Functional Functional ±4 mm Functional |
| Miscellaneous accessories Detection Latching Position Functionality | Functionality Functionality Constancy Functionality | Monthly Monthly Monthly Monthly | Half-yearly Half-yearly Half-yearly Half-yearly | Functional Functional ±2 mm Functional | Functional Functional ±4 mm Functional |

Table II.	Beam-related checks for conventional linear accelerators for 3DCRT

| PARAMETER | CRITERION | RECOMMENDED FREQUENCY | MINIMAL FREQUENCY | TOLERANCE LEVEL | IMMEDIATE ACTION LEVEL |
| --- | --- | --- | --- | --- | --- |
| Beam geometry |  |  |  |  |  |
| Radiation isocenter Gantry rotation axis Collimator rotation axis | Accuracy Accuracy | Half-yearly Half-yearly | Yearly Yearly | ±1 mm ±1 mm | ±2 mm ±2 mm |
| Light and radiation field coincidence Reference SSD and clinical range of field size Clinical range of SSD Gantry angle dependence Field size < 10 x 10 cm² Field size > 10 x 10cm² | Accuracy Accuracy | Monthly Half-yearly Half-yearly | Half-yearly Yearly Yearly | ±1 mm ±1% of field size (max. ±2 mm) | ±2 mm ±2% of field size (max. ±4 mm) |

| Opposing fields radiation field (photon beams) 90°-270° or 0°-180° alternating, photon energy (if more then 1) alternating | Accuracy | Monthly | Half-yearly | ±2 mm | ±2 mm |
| --- | --- | --- | --- | --- | --- |
| Photon beam Flatness1 Symmetry Constancy of Flatness and Symmetry* Gantry angle dependence of 18.a, 18.b and 18.c * Preferably the complete profile should be compared to the reference profile | Accuracy Accuracy Constancy Accuracy | Three-Monthly Three-Monthly Three-monthly Half-yearly | Half-yearly Half-yearly Half-yearly Yearly | 3% 2% 2% idem | 6% 4% 2% idem |
| e. Wedge profiles (relative local deviation over 80 % of the central region): |
| Mechanical Motorised Dynamic | Constancy Constancy Constancy | Yearly Yearly Monthly | Yearly Yearly Half-Yearly | 3% 3% 3% | 3% 3% 3% |
| Electron beam Flatness1 Symmetry1 Dose profiles at clinical SSD and depth range Gantry angle dependence of 19.a, 19.b and 19.c | Accuracy Accuracy Constancy Accuracy | Three-Monthly Three-Monthly Three-monthly Half-yearly | Half-yearly Half-yearly Half-yearly Yearly | 3% 3% 2% idem | 6% 6% 2% idem |
| Remark 1: if electronic adjustment of the beam is not possible, then tolerance/action can be set to 3%/6% respectively |
| Dosimetry system |  |  |  |  |  |
| Performance Gantry angle output dependence Dose rate output dependence Dosimetry system linearity Dosimetry system reproducibility Dosimetry system background count* (* if applicable) | Constancy Constancy Accuracy Accuracy Accuracy | Half-yearly Half-yearly Half-yearly Half-yearly Monthly | Yearly Yearly Yearly Yearly Three-monthy | ±2% ±0.5% 1% 0.5% 1 MU/min | ±4% ±1% 2% 1% 2 MU/min |
| Safety Time interlock Backup MU system | Functionality Functionality | Half-yearly Half-yearly | Yearly Yearly | Functional Functional | Functional Functional |
| Photon beam central axis dosimetry |  |  |  |  |  |
| Beam output at reference deptht Quick Constancy check Dosimetry with field instrument and calibration of quick check Reference dosimetry and calibration of field instrument | Constancy Constancy Accuracy | Daily Monthly Three-monthly | Weekly Three-monthly Half-yearly | ±3% ±2% ±1% of reference dose | ±5% ±3% ±2% |
| Depth dose Quality index (point measurement) Depth dose curve: QI/Depth of maximum dose | Accuracy Constancy | Monthly Three-monthly | Half-yearly Half-yearly | QI ±2% ±2%/2 mm | QI ± 3% ±2% /4 mm |
| Output factors Field size over clinical range Accessory transmission factors Wedge factors (Dopen/Dwedge) Mechanical Motorised Dynamic/Virtual Gantry angle dependence of 24.a, 24.b and 24.c | Constancy Constancy Constancy Constancy Constancy Accuracy | Half-yearly Half-yearly Monthly Monthly Weekly Half-yearly | Yearly Yearly Half-Yearly Half-yearly Monthly Yearly | ±2% ±2% ±2% ±2% ±2% idem | ±4% ±4% ±4% ±4% ±4% idem |
| Electron beam central axis dosimetry |  |  |  |  |  |
| Beam output at reference deptht Quick Constancy check Dosimetry with field instrument and calibration of quick check Reference dosimetry and calibration of field instrument | Constancy Constancy Accuracy | Daily Monthly Three-monthly | Weekly Three-monthly Half-yearly | ±3% ±2% ±1.5% of reference dose | ±5% ±3% ±3% |
| Depth dose Energy constancy check (point measurement) Depth dose curve Depth of maximum dose X-ray contamination | Accuracy Constancy Constancy Constancy | Monthly Three-monthly Three-monthly Three-monthly | Half-yearly Yearly Yearly Yearly |  E0 ± 0.5 MeV ±2%/2mm ±2 mm ±1% |  E0 ± 1.0 MeV ±4% ? ±4 mm ±2% |
| Output factors Field size over clinical range for different applicators Gantry angle dependence of 27.a | Constancy Constancy | Half-yearly Yearly | Yearly Yearly | ±2% ±2% | ±4% ±4% |

Table III.	Machine integrity and safety

| PARAMETER | CRITERION | RECOMMENDED FREQUENCY | MINIMAL FREQUENCY | TOLERANCE LEVEL | IMMEDIATE ACTION LEVEL |
| --- | --- | --- | --- | --- | --- |
| Room entrance interlock / last person out system | Functionality | Monthly | Yearly | Functional | Functional |
| Manual door opening | Functionality | Half-yearly | Yearly | Functional | Functional |
| Patient monitoring devices | Functionality | Monthly | Yearly | Functional | Functional |
| Beam status indicators | Functionality | Monthly | Yearly | Functional | Functional |
| Emergency off switches | Functionality | Monthly | Yearly | Functional | Functional |
| Collision detection system | Functionality | Monthly | Half-yearly | Functional | Functional |
| Motion enable system/ Motion stop switches | Functionality | Monthly | Half-yearly | Functional | Functional |
| Couch brakes | Functionality | Monthly | Half-yearly | Functional | Functional |
| Accessory interlocks | Functionality | Monthly | Half-yearly | Functional | Functional |

* A number of the parameters from table III can be supposed to be inherently checked by the operators/treatment technicians. Instructions should exist that whenever problems are experienced a responsible person is notified and if needed corrective action is organized. However a formal testing of these parameters is recommended at the frequency as specified.